## BSC-1x04IIA2(CE)
Model/Item | BSC-1004IIA2 | BSC-1304IIA2 | BSC-1604IIA2 | BSC-1804IIA2
Nominal Size | 1meter(3') | 1.3meter(4') | 1.6meter(5') | 1.8meter(6')
Filtration Efficiency | ULPA:≥99.9995%, @0.12μm
Downflow Velocity | 0.35 m/s(68.88fpm) |  |  | 0.33 m/s(64.94fpm)
Inflow Velocity | 0.55 m/s(108.24fpm) |  |  | 0.53 m/s(104.30fpm)
Noise | 58~65 dB(A)
Vibration(μm) | ≤5 μm(center of tabletop)
Power Supply | AC220V，1φ，50Hz
Opening Height(mm) | 200
Operator Protection | A. Total colony in impaction sampler ≤10CFU./time
B. Total colony in slot sampler ≤5CFU./time
Product Protection | Total colony in culture dish ≤5CFU./time
Cross Contamination | Total colony in culture dish ≤2CFU./time
Max Consumption(with spare socket) | 1.65 kVA | 1.75 kVA | 1.85 kVA
Rated Power(without spare socket) | 0.33 kVA | 0.43 kVA | 0.53 kVA
Weight | 260 kg | 300 kg | 360 kg | 400 kg
Work Dimension | 1004*630*630 mm | 1304*630*630 mm | 1604*630*630 mm | 1804*630*630 mm
(W1XD1XH1) | (39.5''x24.8''x24.8'') | (51.3''x24.8''x24.8'') | (63.1''x24.8''x24.8'') | (71.0''x24.8''x24.8'')
Overall Dimension | 1200*795*2050 mm | 1500*795*2050 mm | 1800*795*2050 mm | 2000*795*2050 mm
(WXDXH) | (47.2''x31.3''x80.7'') | (59.1''x31.3''x80.7'') | (70.9''x31.3''x80.7'') | (78.7''x31.3''x80.7'')
Size and Qty. of Light | LED 24.5 w x ① | LED 31 w x ① | LED 36 w x ①
Size and Qty. of UV Light | 18 w x ① | 30 w x ① | 36 w x ①
Lumin. | ≥900 Lx
Cabinet Materials | High-grade steel and lacquered in ivory
Working Area Materials | SS304 fully finished
Air Direction | Top out
Plug Style | Chinese Type(AC250V/10A)
Qty. of Spare Socket | Chinese Type(AC250V/3A) 2 units

## BSC-1301_1302ⅡA2(NSF)
Model/Item | BSC-1301ⅡA2 | BSC-1302ⅡA2
Nominal Size | 1.3meter(4')
Filtration Efficiency | HEPA: ≥99.995%, @0.3μm/ULPA:≥99.9995%, @0.12μm
Downflow Velocity | 0.35 m/s(68.88fpm)
Inflow Velocity | 0.54 m/s(106.27fpm)
Noise | ≤67 dB(A)
Vibration(μm) | ≤5 μm(center of tabletop)
Power Supply | AC115V，1φ，60Hz | AC230V，1φ，50/60Hz
Opening Height(mm) | 250
Operator Protection | A. Total colony in impaction sampler ≤10CFU./time
B. Total colony in slot sampler ≤5CFU./time
Product Protection | Total colony in culture dish ≤5CFU./time
Cross Contamination | Total colony in culture dish ≤2CFU./time
Max Consumption(with spare socket) | 1.5 kVA | /
Rated Power(without spare socket) | 1 kVA | 0.46 kVA
Weight | 300 kg
Work Dimension | 1270*600*620 mm
(W1XD1XH1) | (50''x23.6''x24.4'')
Overall Dimension | 1500*790*2050 mm
(WXDXH) | (59.1''x31.1''x80.7'')
Size and Qty. of Light | LED 24.5 w x ②
Size and Qty. of UV Light | 15 w x ②
Lumin. | ≥750 Lx
Cabinet Materials | High-grade steel and lacquered in ivory
Working Area Materials | SS304 fully finished
Air Direction | Top out
Plug Style | USA Type(AC125V/10A) | /
Qty. of Spare Socket | USA Type(AC125V/2A) 2 units | /